Technical Documentation

PulseCare BP-200 Home Blood Pressure Monitor

Document no. TD-PC200-001, Revision C, 14 September 2026

Note: MedNova GmbH and the PulseCare BP-200 are invented for a course project. This document is fictional and must not be used for any real product.

# 1. Manufacturer information

Manufacturer: MedNova GmbH, Rheinallee 48, 50668 Köln, Germany.

Contact for regulatory matters: Regulatory Affairs Department, regulatory@mednova-example.de, +49 221 555 0100.

Single registration number (SRN): DE-MF-000099999 (fictional).

# 2. Device description

The PulseCare BP-200 is a battery-powered, automatic, non-invasive blood pressure monitor for use on the upper arm. It measures systolic and diastolic blood pressure and pulse rate using the oscillometric method.

Main components: main unit with LCD display, inflatable cuff (22–42 cm arm circumference), air tube, rechargeable lithium-ion battery (3.7 V, 1800 mAh), USB-C charging port and Bluetooth Low Energy module for transfer of readings to the PulseCare mobile app.

The device stores up to 120 readings for each of two users. Software version: firmware 2.4.1.

Classification: Class IIa medical device according to Annex VIII, Rule 10 of Regulation (EU) 2017/745 (MDR).

# 3. Intended purpose

The PulseCare BP-200 is intended to measure the systolic and diastolic blood pressure and pulse rate of adults (18 years and older) at home, using a cuff on the upper arm.

Intended users: lay persons (patients) without medical training. Intended environment: home use.

Contraindications: not intended for use on neonates, infants or pregnant women with pre-eclampsia, and not intended for continuous monitoring in hospitals.

# 4. Risk management

Risk management was performed according to ISO 14971:2019. The risk management plan is document RMP-PC200-001. The table below summarises the main identified risks and the control measures applied.

Table 4.1 Risk analysis summary

| Risk ID | Hazard | Possible harm | Control measure |
| --- | --- | --- | --- |
| R-01 | Incorrect blood pressure reading | Wrong treatment decision | Clinical validation per ISO 81060-2; error message if movement is detected during measurement |
| R-02 | Excessive cuff pressure | Bruising, pain in the arm | Hardware pressure limit at 300 mmHg; automatic deflation after 180 seconds |
| R-03 | Battery overheating during charging | Burns to the skin |  |
| R-04 | Electric shock from charger | Electric shock | Use of certified 5 V medical-grade USB adapter; device cannot be used while charging |
| R-05 | Loss of data transfer via Bluetooth | Missing readings in app | Readings stored on device; user notified if transfer fails |

All residual risks were evaluated as acceptable in relation to the benefits of home blood pressure monitoring. The overall residual risk is acceptable.

# 5. Verification and test results

Electrical safety and performance tests were carried out by NordCert Prüflabor GmbH (test report TR-2026-0417). The following standards were applied: IEC 60601-1:2005+A2:2020 (general safety), IEC 60601-1-2:2014 (EMC), IEC 60601-1-11:2015 (home healthcare environment), IEC 80601-2-30:2018 (automated non-invasive sphygmomanometers) and ISO 81060-2:2018 (clinical validation).

Table 5.1 Summary of test results

| Test | Standard / clause | Limit | Measured value |
| --- | --- | --- | --- |
| Patient leakage current (normal condition) | IEC 60601-1, 8.7 | ≤ 100 µA | 45 µA |
| Cuff pressure accuracy | IEC 80601-2-30, 201.12.1.102 | ± 3 mmHg | ± 2 mmHg |
| Maximum cuff pressure | IEC 80601-2-30, 201.12.4.103 | ≤ 300 mmHg | 292 mmHg |
| Surface temperature of battery compartment during charging | IEC 60601-1, 11.1 | ≤ 60 °C | 72 °C |
| Drop test (free fall, 1 m) | IEC 60601-1-11, 10.1.3 | No damage affecting safety | No damage |
| Ingress protection | IEC 60529 | IP22 | IP22 |

Clinical validation according to ISO 81060-2 was performed with 85 subjects. Mean difference: 1.2 mmHg (systolic) and 0.8 mmHg (diastolic), within the required ± 5 mmHg.

# 6. Labelling and instructions for use

To be completed.

# 7. Software

The firmware was developed according to IEC 62304:2006+A1:2015. Software safety class: Class B. The software requirements, architecture and verification are documented in SW-PC200-002.

# 8. Biocompatibility

The cuff material in contact with skin (nylon outer layer, TPU bladder) was evaluated according to ISO 10993-1. Tests for cytotoxicity, sensitisation and irritation passed.

# 9. Post-market surveillance

MedNova GmbH maintains a post-market surveillance plan (PMS-PC200-001). Complaints are reviewed monthly. A periodic safety update report (PSUR) is prepared every two years.

# 10. Document approval

Table 10.1 Revision history

| Revision | Date | Change | Approved by |
| --- | --- | --- | --- |
| A | 12 March 2026 | First release | Dr. Lena Hoffmann, Head of Quality |
| B | 30 June 2026 | Updated risk analysis | Dr. Lena Hoffmann, Head of Quality |
| C | 14 September 2026 | Added test results | Dr. Lena Hoffmann, Head of Quality |