NordCert Prüflabor GmbH

Test Report TR-2026-0417

Electrical safety testing according to IEC 60601-1

Note: NordCert Prüflabor GmbH, MedNova GmbH and the PulseCare BP-200 are invented for a course project. This report is fictional.

# 1. General information

| Item | Details |
| --- | --- |
| Testing laboratory | NordCert Prüflabor GmbH, Hafenstraße 9, 20457 Hamburg, Germany |
| Client | MedNova GmbH, Rheinallee 48, 50668 Köln, Germany |
| Product | Home blood pressure monitor |
| Model | PulseCare BP-200 |
| Serial number of sample | PC200-2026-00117 |
| Date of receipt of sample | 4 August 2026 |
| Testing period | 11 August 2026 to 22 August 2026 |
| Date of issue | 29 August 2026 |

# 2. Test standards

The sample was tested according to IEC 60601-1:2005+A2:2020 (Medical electrical equipment – Part 1: General requirements for basic safety and essential performance) and IEC 60601-1-2:2014 (electromagnetic disturbances).

# 3. Test plan

The following tests were agreed with the client in the test plan TP-2026-0417: patient leakage current, dielectric strength, surface temperatures, mechanical strength, and EMC immunity.

# 4. Test results

Table 4.1 Test results

| Test | Clause | Requirement | Measured value | Verdict |
| --- | --- | --- | --- | --- |
| Patient leakage current | 8.7 | ≤ 100 µA | 45 µA | Pass |
| Dielectric strength | 8.8.3 | No breakdown at 1500 V AC | No breakdown | Pass |
| Surface temperature, battery compartment | 11.1 | ≤ 60 °C | 72 °C | Fail |
| Mechanical strength (push test 250 N) | 15.3.2 | No damage | No damage | Pass |
| EMC immunity | IEC 60601-1-2 | Essential performance maintained | Not performed: EMC chamber not available | — |

# 5. Conclusion

The sample does not fully comply with IEC 60601-1. The surface temperature of the battery compartment exceeded the limit during charging. The EMC immunity test is still open.

# 6. Signatures

Tested by: Markus Brandt, Test Engineer, 29 August 2026

Reviewed and approved by: ______________________   Date: __________